| Laser Type: | 980 nm Diode laser |
| --- | --- |
| Wavelength (nm): | 980 nm |
| Beam Diameter (mm): | 1 |
| Beam Diameter (mm): | 1 |
| Mode: | continuous-wave |
| Pulse Duration (sec): |  |
| Pulse Frequency (Hz): |  |
| Average Joules/Pulse (J): |  |
| Maximum Joules/Pulse: |  |
| Average Power (watts): | 0.515 |
| Maximum Power (watts): | 0.515 |
| Laser Type #1: | 980 nm laser |
| Operation Power (W or J): | 515 mW |
| Wavelength (nm): | 980 |
| Wavelength Attenuated: | N/A |
| Optical Density Required: | 5+ |
| Eyewear (Mfg/model): | Thorlabs - LG11 |
| When use is required: | Always |
| Additional Information: | Because the laser is invisible OD 5+ is required at all times. Also, cover the beam path whenever possible |
| Precaution notes: | The laser enters the microscope from below so make sure to not look down at the microscope while measuring. Cover the sample with a beam block whenever possible. |
| Protection notes: | There is compressed gas for the optical table to float. Be aware of that hazard while working. |
| Operating notes: | General Operating Procedures: - Put on laser safety goggles - Turn on the lab laser safety light - Close laser lab entrance curtains Alignment: - Ensure that laser is firmly mounted to posts. - Check that path of beam ends suitable beam dump or a fiber coupler - Turn key on laser control box - Detector card in front of laser should show a green spot. - Flip this up when ready to align. - Align at low powers using detector card and/or IR viewer. Note that laser can burn alignment card if tightly focused or held in beam path for too long. Start-up: - Laser takes ~ 5 minutes to stabilize after initial turn-on. - When not actively taking measurements, turn off the key switch to stop lasing Shutdown: - Turn off laser controller. - Turn off laser light. |