Subject: Laser Standard Operating Procedure

| Laser Type: | CO2 |
| --- | --- |
| Wavelength (nm): | 10600 |
| Beam Diameter (mm): | 1 |
| Beam Diameter (mm): | 1 |
| Mode: | continuous |
| Pulse Duration (sec): |  |
| Pulse Frequency (Hz): |  |
| Average Joules/Pulse (J): |  |
| Maximum Joules/Pulse: |  |
| Average Power (watts): | .78 |
| Maximum Power (watts): | .78 |
| Laser Type #1: | 10600 nm laser |
| Operation Power (W or J): | 780 mW |
| Wavelength (nm): | 10600 |
| Wavelength Attenuated: | N/A |
| Optical Density Required: | at least 5 |
| Eyewear (Mfg/model): | Glendale – white frames |
| When use is required: | Always |
| Additional Information: | Eyeware will be reviewed annually and unfit/damaged eyeware will be discarded |
| Precaution notes: | - Verify that the beam block and any necessary optics are in place. For instance, the laser head is directed towards a beam dump (non-reflective, non-flammable dark surface). |
| Protection notes: | - There is compressed gas for the optical table to float. Be aware of that hazard while working. |
| Operating notes: | General Operating Procedures: - Obtain key to operate the laser from the lab laser safety supervisor. Never give the key to a third party without permission - Clear the lab of non-essential personnel and anyone not equipped with proper safety gear. Make sure only individuals with the appropriate safety training are in the lab. - Close the entry door to the laser lab spaces then close the curtain of the room where the laser is going to be operated. Switch on the laser warning sign outside the laboratory. - Place the “LASER IN USE!! – KNOCK LOUDLY FOR ACCESS” sign over the keypad. Verify that the lock is working. Alignment & Calibration: - See user manual for details on alignment and calibration - Only conduct alignment or calibration procedure if you have been properly trained Start-up & Operation: - Verify all persons in lab wear eyewear with appropriate optical density and wavelength. - Place the sample at the sample stage for experiments. - Turn on the laser power supply with key switch and wait for the display to load. - Open shutter when ready and adjust laser power when needed for collecting high-quality data. - Keep the room curtain closed when collecting data over a long period and anytime the laser is energized and operating with accessible laser radiation. - Close the laser shutter while the laser is not in use. Shutdown: - Close the laser shutter. Reduce laser current to minimum operational value. - Turn off laser and switch the key to the standby position. - Remove the “Laser in Use” sign from the entry door. - Clear up all clutter, samples, and tools, and reset any changes made to the laser set up. - After use, return the key to the lab laser safety officer immediately |